## Comments
INTERIM STANDARD. ELV1M4, SECTIONS 1.7.1 AND 1.7.2 (CALIBRATION). APRIL 2015. RESPONSE TO COMMENTS
Vote tally: Abstain 2; Affirmative 27; Affirmative with Comment 5; Negative with Comment 13
Disclaimer: The NELAC Institute (TNI) accepts no liability for the content of any comment made on a standard. Any views or opinions on a standard are solely those of the commenter and do not necessarily reflect those of TNI.
| Vote | Section | Comment | Disposition | Change/Reasoning
| Affirmative with comment | General | This is an improvement in the standard. The CEC did an excellent job in developing this standard and considering stateholder comments. | N/A | Thank you for your comment
| Affirmative with comment | General | I strongly advocate adding language to better clarify what methods do and don't fall into the category of being covered by this section. For instance, conductivity is traditionally, and by method, calibrated by determining a cell constant, essentially a one point calibration forced through zero. This would seem to be included as being covered by this section as it passes the '...methods that use calibration models including but not limited to average response factor or linear or quadratic regression...' criteria. I don't believe this was the intention of the committee as there are requirements in the standard that can't be met, or don't make sense, for conductivity analysis - such as, 1.7.1.1, f) and g) and k) ii). Adding a statement that single point calibrations (such as threshold testing and conductivity measurements) are excluded from these requirements might help to clarify. (Although I see that threshold testing is mentioned in part 1.7.1.1 e) which implies that the other parts of 1.7.1.1 apply to threshold testing which, again, I don't believe was the intention.) | Persuasive | We discussed at length, but were ultimatly unable to agree on suitable language to address this point. This comment will be placed on hold and considered during the next revision cycle of the standard.
| Affirmative with comment | Multiple | Mostly editorial changes and correction of passive voice language. See the uploaded file, which is a Word version of the standard with my changes tracked and followed by the rationale in bold red font. Thanks! | Persuasive | Thankyou for these suggestions. We have incorporated most into the standard language.
| Negative with comment | All | I appreciate the Committee's work on the interim standard -- especially their efforts to incorporate changes based on the Working Draft Standard. The Interim Standard is a significant improvement on the WDS. I am somewhat skeptical that the additional work required by the Interim Standard will provide enough significant practical improvements in data quality to offset the additional laboratory requirements. Therefore, I am voting negative. Since the instructions indicated a comment was required for my vote to be considered, I have entered the above comments. | Non-Persuasive | Thankyou for your comment. We believe that the changes to the standard will result in substantial data quality improvements, but admit it is a matter or opinion.
| Negative with comment | 1.7.1.1 (d) (ii) | I believe that for tests with a large number of analytes the requirement to always remove an entire calibration level(all analytes) from the interior of a calibration curve is excessive. To me it makes more sense for you to remove all potentially effected analytes. To better clarify, an example that we have at our lab would be our standard 8260 list which contains 105 analytes. A calibration curve contains 11 points and each point is spiked with five standards. That is 55 individual spikes per curve. It does happen where only one of the five standards was spiked incorrectly or a spiking error occurred in one of the calibration points. In this case currently we would document that all compounds from standard "X" were removed due to spiking error. There is absolutely nothing wrong with all the other compounds in the calibration level, and when an incident like this occurs it is clearly identifiable. In my opinion the above stated practice with the proper documentation should not be excluded from the standard. We should not be required to remove data points that are clearly acceptable due to the failure of one spike mix or any other clearly identifiable cause for specific compound/compound group failure. An error of this nature is bound to happen from time to time due to the high number of analytes and individual spikes required. To take it a step further I believe similar logic could be used here as is used for the marginal exceedence rule. A certain number of failures should be understood if not expected when you are working with this number of analytes and standards. In the case of the 8260 curve listed above the instrument would be generating 1155 individual data points. It is completely understandable that everything will not function perfectly for each data point. |  | Comment withdrawn
| Negative with comment | 1.7.1.1 (d) (ii) | I believe that for tests with a large number of analytes the requirement to always remove an entire calibration level(all analytes) from the interior of a calibration curve is excessive. To me it makes more sense for you to remove all potentially effected analytes. To better clarify, an example that we have at our lab would be our standard 8260 list which contains 105 analytes. A calibration curve contains 11 points and each point is spiked with five standards. That is 55 individual spikes per curve. It does happen where only one of the five standards was spiked incorrectly or a spiking error occurred in one of the calibration points. In this case currently we would document that all compounds from standard "X" were removed due to spiking error. There is absolutely nothing wrong with all the other compounds in the calibration level, and when an incident like this occurs it is clearly identifiable. In my opinion the above stated practice with the proper documentation should not be excluded from the standard. We should not be required to remove data points that are clearly acceptable due to the failure of one spike mix or any other clearly identifiable cause for specific compound/compound group failure. An error of this nature is bound to happen from time to time due to the high number of analytes and individual spikes required. To take it a step further I believe similar logic could be used here as is used for the marginal exceedence rule. A certain number of failures should be understood if not expected when you are working with this number of analytes and standards. In the case of the 8260 curve listed above the instrument would be generating 1155 individual data points. It is completely understandable that everything will not function perfectly for each data point. | Non-Persuasive | We discussed these points at length, but ultimately decided that the change that you request would likely be considered negatively by part of the TNI community. We will continue to consider improvements to the language for the next version of the standard
| Negative with comment | 1.7.1.1 (d) | As an assessor, we must have written procedures when removal or replacement of points occur. This should not have been removed - having this allows assessors to clearly write this as a finding. | Persuasive | Added language to 1.7.1.1 d
| Negative with comment | 1.7.1.1 (d) | This section has been revised to remove the requirement for a laboratory to have a written procedure to address removal/replacement of calibration standards. Even though the standard is more clear regarding removal/replacement, a written procedure is still necessary so that the laboratory can define its actual procedures and address how and where the required documentation will take place. The procedure is also essential to staff training and consistent application of the use of removal/replacement. | Persuasive | Added language to 1.7.1.1 d
| Negative with comment | 1.7.1.1 (e) | It is our opinion that it is not necessary to increase the minimum number of calibration standards. Since the degrees of freedom portion of the table is deleted there is no need for Note (b). Delete Note (b). | Non-Persuasive | The standard, as written, now has a statistical basis for the minimum number of standards required for each type of calibration. This requirement is also consistent with the requirements specified in the current EPA SW-846 methods and updates to the EPA 600 series methods for the minimum number of initial calibration standards for the different calibration types. Note b will be reworded as follows: “Fewer calibration standards (resulting in less than 3 statistical degrees of freedom) may be used……..”
| Negative with comment | 1.7.1.1 (e) | It is our opinion that it is not necessary to increase the minimum number of calibration standards. Since the degrees of freedom portion of the table is deleted there is no need for Note (b). Delete Note (b). | Non-Persuasive | The standard, as written, now has a statistical basis for the minimum number of standards required for each type of calibration. This requirement is also consistent with the requirements specified in the current EPA SW-846 methods and updates to the EPA 600 series methods for the minimum number of initial calibration standards for the different calibration types. Note b will be reworded as follows: “Fewer calibration standards (resulting in less than 3 statistical degrees of freedom) may be used……..”
| Negative with comment | 1.7.1.1 (e) | Do not increase the minimum number of calibration standards, this is unnecessary. Delete Note (b). Since the degrees of freedom portion of the table is deleted there is no need for Note (b). | Non-Persuasive | The standard, as written, now has a statistical basis for the minimum number of standards required for each type of calibration. This requirement is also consistent with the requirements specified in the current EPA SW-846 methods and updates to the EPA 600 series methods for the minimum number of initial calibration standards for the different calibration types. Note b will be reworded as follows: “Fewer calibration standards (resulting in less than 3 statistical degrees of freedom) may be used……..”
| Abstain | 1.7.1.1 (e) | Table requires a minimum of 4 calibration standards for Average Response calibration curve fit; however 600 series methods including 624 and 625 allow for Average Response calibration fits, but only require a minimum of 3 calibration standards. This would provide conflict between the method and the TNI standard. | Non-Persuasive | The standard, as written, now has a statistical basis for the minimum number of standards required for each type of calibration. This requirement is also consistent with the requirements specified in the current EPA SW-846 methods and updates to the EPA 600 series methods for the minimum number of initial calibration standards for the different calibration types. Note b will be reworded as follows: “Fewer calibration standards (resulting in less than 3 statistical degrees of freedom) may be used……..”
| Negative with comment | 1.7.1.1 (e) | The 600 series methods require 3 stds for linear regression but now this requires 4 points. What is the basis for this? This will increase work with little benefit in quality. | Non-Persuasive | The standard, as written, now has a statistical basis for the minimum number of standards required for each type of calibration. This requirement is also consistent with the requirements specified in the current EPA SW-846 methods and updates to the EPA 600 series methods for the minimum number of initial calibration standards for the different calibration types. Note b will be reworded as follows: “Fewer calibration standards (resulting in less than 3 statistical degrees of freedom) may be used……..”
| Negative with comment | 1.7.1.1 (e) | Calibration table. The minimum number of standards for Average Response changed from 3 to 4, and is still in conflict with the minimum as represented in the comment I submitted in April 2014. See original comment below: The chart requires a minimum of 4 calibration standards for linear regression, but this goes against multiple EPA methods (i.e. EPA 625 requires a minimum of 3 calibration standards, but allows for linear regression fits where RSD does not meet acceptance criteria and the same would hold true for multiple 600 series methods). The required number in this chart for Quadratic Fit is also in conflict with SW846 EPA 8000C that states in Section 11.5.3.1, " . . .a quadratic (second order) model requires six standards, . . ." Five standards for a quadratic fit would not adequately model a curve with an inflection point and may cause gross errors in the quantitation of target analytes at a mid-level concentration. | Non-Persuasive | The standard, as written, now has a statistical basis for the minimum number of standards required for each type of calibration. This requirement is also consistent with the requirements specified in the current EPA SW-846 methods and updates to the EPA 600 series methods for the minimum number of initial calibration standards for the different calibration types. Note b will be reworded as follows: “Fewer calibration standards (resulting in less than 3 statistical degrees of freedom) may be used……..”
| Negative with comment | 1.7.1.1 (g) | Add an allowance to report results that are within +10% of the highest calibration level without qualification. For example, if the high standard is 300 and you obtain a result of 301 you would need to qualify and/or dilute the sample before reporting. Dilution would introduce more error than reporting at the original 301 concentration. | Non-Persuasive | This is a request for new item in the standard that cannot be considered at the interim standard stage. The request will be place on hold and considered during the next revisiion cycle of the standard.
| Negative with comment | 1.7.1.1 (j) | It is our opinion that calculating a %RE or %RSE as a means of assessing the acceptability of calibration curves would not be cost effective and as an alternative, we suggest using similar language as used in UCMR3 Methods. As an example Method 524.3, rev 1.0, June 2009. This procedure is easily implemented offers EPA acceptance, and consistency with acceptance criteria across all laboratories. Method 524.3, rev 1.0 June 2009 – 10.1.10 CALIBRATION ACCEPTANCE CRITERIA – The initial calibration is validated by calculating the concentration of the analytes for each of the analyses used to generate the calibration curve by use of the regression equations. Calibration points that are ≤MRL must calculate to be within +50% of their true value. All other calibration points must calculate to be within +30% of their true value. If these criteria cannot be met, the analyst will have difficulty meeting ongoing QC criteria. In this case, corrective action is recommended such as reanalyzing the calibration standards, restricting the range of calibration, or performing instrument maintenance. | Non-Persuasive | The committee concurs that the calibration acceptance criteria as specified in UCMR3 methods is reasonable and in harmony with the %RE approach as specified in this section of the standard. However, this approach is too narrow to satisfy all calibration types. The standard as written, allows the use of %RE or %RSE, thus providing additional flexibility to the user.
| Negative with comment | 1.7.1.1 (j) | It is our opinion that calculating a %RE or %RSE as a means of assessing the acceptability of calibration curves would not be cost effective and as an alternative, we suggest using similar language as used in UCMR3 Methods. As an example Method 524.3, rev 1.0, June 2009, section 10.1.10 Calibration Acceptance Criteria. This procedure is easily implemented offers EPA acceptance, and consistency with acceptance criteria across all laboratories. Read back calibration response off the curve and compare recovery results to the method calibration verification limits for all points except those that are at or below the MRL, assess at 50% recovery. | Non-Persuasive | The committee concurs that the calibration acceptance criteria as specified in UCMR3 methods is reasonable and in harmony with the %RE approach as specified in this section of the standard. However, this approach is too narrow to satisfy all calibration types. The standard as written, allows the use of %RE or %RSE, thus providing additional flexibility to the user.
| Negative with comment | 1.7.1.1 (j) | Calculating a %RE or %RSE as a means of assessing the acceptability of calibration curves would not be cost effective and as an alternative, we suggest using similar language as used in UCMR3 Methods. As an example Method 524.3, rev 1.0, June 2009, see below. This procedure is easily implemented, offers EPA acceptance, and consistency with acceptance criteria across all laboratories. Method 524.3, rev 1.0 June 2009 – 10.1.10 CALIBRATION ACCEPTANCE CRITERIA – The initial calibration is validated by calculating the concentration of the analytes for each of the analyses used to generate the calibration curve by use of the regression equations. Calibration points that are ≤MRL must calculate to be within +50% of their true value. All other calibration points must calculate to be within +30% of their true value. If these criteria cannot be met, the analyst will have difficulty meeting ongoing QC criteria. In this case, corrective action is recommended such as reanalyzing the calibration standards, restricting the range of calibration, or performing instrument maintenance. | Non-Persuasive | The committee concurs that the calibration acceptance criteria as specified in UCMR3 methods is reasonable and in harmony with the %RE approach as specified in this section of the standard. However, this approach is too narrow to satisfy all calibration types. The standard as written, allows the use of %RE or %RSE, thus providing additional flexibility to the user.
| Negative with comment | 1.7.1.1 (j) | "If there is no specification for RSE or RSD in the method, then the RSE shall be specified in the laboratory SOP." A majority of the inorganic methods do not have criteria for RSD or RSE. Without any guidance being provided by either the published method or regulatory guidelines for either RSD or RSE, how can an assessor consistently audit laboratories to the calibration section of the standard. Each lab could have very different acceptance criteria for either parameter (RSE or RSD) and as long as the laboratory meets it’s defined requirement, an assessor could not cite for inadequacy of one lab’s practice against another. Therefore, the consistency of the evaluation process is compromised due to incomplete guidance from the TNI Standard. "If no criterion for the lowest calibration level is specified in the method, the criterion and the procedure for deriving the criterion shall be specified in the laboratory SOP." As stated in above comment how can an assessor consistently evaluate laboratories if each laboratory chooses their own criteria instead of using criteria with guidance from regulatory agencies or the TNI Standard. | Non-Persuasive | The committee recognizes that for many methods %RE and %RSE criteria do not currently exist, but we feel that adding a measure of relative uncertanty to the calibration is important for improving the quality of data generated, especially at the low end of the calibration curve. We feel that reasonable acceptance criteria will be developed by laboratories over time and at some point in the future more specific industry wide perfromance criteria can be defined. This is not the only area in TNI standards where the laboratory is left to establish their own acceptance criteria. As newer methods are issused over time, more and more criteria for calibration uncertainty will be established.
| Abstain | 1.7.1.1 (j) | Criteria for %RE and %RSE are not included in any of the published methods historically used for environmental analysis. I understand that it has been included in the updated version of EPA 8000; however not all states have moved to the updated versions of the SW-846 methods. This could cause a lapse in assessor understanding and application of the concepts given. Drift and %Difference are normally listed for daily calibration verification use, but appear to be much the same as the equations given in the proposed standard for %RE. Why confuse the nomenclature by introducing different terminology for these concepts when they are clear in the published methods. In addition, since no criteria is given for acceptance of these concepts in the draft standard, using the acceptance criteria in the methods along with the terminology in the methods seems more appropriate for the evaluation of the mid-level standard. The inclusion of the low level standard is a variation from method protocols and no supporting data is available that demonstrates that assessing the low level standard to the same criteria as a mid-level standard would be practical in real world applications. With no acceptance criteria given in the standard, you are allowing each laboratory to determine a criteria for evaluation would provide inconsistency of data between laboratories and would force assessors to develop independent acceptance criteria for each regulatory program that would then remove consistency in the evaluation process that was the initial and ultimate goal of the TNI program. | Non-Persuasive | The committee recognizes that for many methods %RE and %RSE criteria do not currently exist, but we feel that adding a measure of relative uncertanty to the calibration is important for improving the quality of data generated, especially at the low end of the calibration curve. The concepts of %Difference and "Drift" as used to assess the CCV are the same mathmatically and also very usefull. We have added a definition for Relative Error which includes clarification as to how %RE relates to % Differnce and Drift to prevent confussion. While evaluating the uncertainty of a calibration at the mid level is a good place to begin, it is not sufficent to control error at the low end of the calibration, which is where most error occurs even with good corelation coefficients. We feel that reasonable acceptance criteria will be developed by laboratories over time and at some point in the future more specific industry wide perfromance criteria can be defined, but for now we prefer to maintian some flexibility in the standard. This is not the only area in the TNI standards where the laboratory is left to establish their own acceptance criteria.
| Negative with comment | 1.7.1.1 (j) | 1.7.1.1 Initial Calibration criteria for the acceptance of an initial calibration shall be established (e.g., correlation) coefficient or relative standard deviation). Ii b) b. measurement of the relative Standard Error (%RSE) Relative Standard Error is calculated using the following equation: This bolded statement in relative error would not relative to a majority but to a limited number of methods. Few, if any, approved methods for inorganic and metals analyses have criteria for RSD. With no criteria, the concern is the potential to misunderstand and misuse RE/RSE. With no previous requirement for RE/RSE, there is no historical performance to establish a basis, therefore placing a regulatory burden on the ABs to determine what is acceptable. Because there is no history for majority of the inorganic and metals methods, there is no way to determine potential benefit, which can defy the purpose of adding new requirements. | Non-Persuasive | The committee recognizes that for many methods %RE and %RSE criteria do not currently exist, but we feel that adding a measure of relative uncertanty to the calibration is important for improving the quality of data generated, especially at the low end of the calibration curve. We feel that reasonable acceptance criteria will be developed by laboratories over time and at some point in the future more specific industry wide perfromance criteria can be defined, but for now we prefer to maintian some flexibility in the standard.
| Negative with comment | 1.7.1.1 (j) | Remove the requirement to calculate %RE or %RSE and allow this to be driven by the referenced methods (i.e. 8260C, 8270D). | Non-Persuasive | The committee concurs that the calibration acceptance criteria as specified in 8260C and 8270D methods is reasonable and in harmony with the approach as specified in this section of the standard. However, this approach is too narrow to satisfy all calibration types and methods. The standard as written, allows the use of %RE or %RSE, thus providing additional flexibility to the user.
| Negative with comment | 1.7.1.1 (j) | Most metals and general chemistry methods do not require an additional RE or RSE for the mid point and low standard in the curve. This will cause a lot of either manual calculation because most data stations doing regression, do not have this built in as an addition. While the calculations are easy, it will require more manual or spread sheet calculations for many organizations. Can the committee clarify what this will buy with respect to quality. If one really wants to evaluate linearity then the low and high should be used not a mid and low point. | Non-Persuasive | The committee feels that adding a measure of relative uncertanty to the calibration is important for improving the quality of data generated, especially at the low end of the calibration curve. Several examples of the value of using %RE and %RSE were provided in Webinars and other presentations. While evaluating the uncertainty of a calibration at the mid level is a good place to begin, it is not sufficent to control error at the low end of the calibration, which is where most error occurs even with good corelation coefficients.
| Negative with comment | 1.7.1.1 (j) | This is a new requirement with no criteria. While laboratories are accustomed to meeting similar requirements for Organic analytical methods, this criteria is new to most inorganic/metals methods. A requirement with no criteria provides no benefit to the standard. Historical practice for calibration curve evaluation relates linearity to a correlation coefficient of >0.995 or a coefficient of determination of >0.99. While this alone does not represent true linearity, adding RE/RSE with no historical basis adds little value. It is possible to have good linearity according to r or r2 and have a bad RE/RSE. But with no historical basis or criteria, who determines what a bad RE/RSE really is? In addition one can have a “bad” RE/RSE at one portion of the curve and a good RE/RSE at a different portion of the curve. Who determines what is acceptable? Is it okay to have a value at the low end of the curve at 40% and a value of 10% at the mid-point of the curve? This would indicate tilting of the curve, but with no acceptance criteria, what determines acceptable and unacceptable? Why pick low and mid points? What is the basis? Seems as though low and high points would be more representative. Curves typically tilt at the mid-point and therefore tend to be more accurate at the mid-point due to the nature of linear “curve-fit”. I agree that there needs to be better linearity evaluation criteria than just evaluating r or r2, but requiring something new with no historical basis is not a proven choice. This would be a good requirement for the next version of the standard and would give the lab community time to generate data to determine criteria and to determine where the two required standards would provide the best information, low, mid or high. | Non-Persuasive | The committee feels that adding a measure of relative uncertanty to the calibration is important for improving the quality of data generated, especially at the low end of the calibration curve. Several examples of the value of using %RE and %RSE were provided in Webinars and other presentations. While evaluating the uncertainty of a calibration at the mid level is a good place to begin, it is not sufficent to control error at the low end of the calibration, which is where most error occurs even with good corelation coefficients. An evaluation of uncertainty at the high end of the curve is certainly allowed in addition to the lower half of the calibration, but the committee felt that was too onerous relative to the added benefit to make it a requirement.
| Affirmative with comment | 1.7.1.1 (m) | The intent of the statement that "...all initial calibrations shall be verified with a standard obtained from a second independently prepared lot or from a second manufacturer" needs clarification. To avoid the possible interpretation by an auditor that someone other than the analyst who prepared the calibration standards must prepare the ICV, I suggest that the clause should be revised to read "...all initial calibrations shall be verified with a standard obtained from a second manufacturer or a separate lot prepared independently by the manufacturer." | Persuasive | Language modified to make initial standards, not ICV, traceable to national standarrds
| Negative with comment | 1.7.1.1 (m) | ISO17025 already requires initial calibration standards that are traceable to a national standard. There is added cost to purchasing these standards. Adding it as an ICV requirement in the TNI Standard is not necessary since the calibration standards are already traceable and the ICV is checked against those. It would be a redundant check with increased cost to the laboratories. | Persuasive | Language modified to make initial standards, not ICV, traceable to national standarrds
| Negative with comment | 1.7.2 (c) and (d) | We often receive client requirements in QAPPs that stipulate the concentration range for the CCV that differ from the </= ½ the high cal std requirement stated in c). There are also some methods and client requirements for a low and high level LCS. Would d) iii. allow using these LCS standards for the CCV when the processes are the same for the two standards? Is it acceptable to use project specific criteria for the CCV levels if it differs from the range stipulated in the TNI standard? | Non-Persuasive | We cannot change this item at this stage since it is potentially controversial and might cause individuals that had voted positively to modify their vote. However we note that it is acceptable to use project or method specific criteria that specify CCV levels.
| Affirmative with comment | 1.7.2 (d) | This whole section is confusing to me. Any clause that contains the word "except" tends to invite further thought or scrutiny. If the Committee would please consider and use the following revisions, this section would read much more clearly (at least to me) on what the laboratory is required to do (changes and additions underlined): d) Instrument continuing calibration verification shall be performed at the beginning and end of each analytical batch, normally with the same standard source as used for the initial calibration, and at the frequency defined in the method except: i. if an internal standard calibration procedure is used, calibration verification needs only to be performed at the beginning of each analytical batch, and at the frequency defined in the method; ii. a calibration verification with a second source standard that passes the continuing calibration verification criteria may be used in place of a continuing calibration verification made with the same initial calibration standard source. iii. a laboratory control sample (LCS) may be used in place of a continuing calibration verification (but not as a replacement for a failing CCV) for methods where the calibration goes through the same process as the LCS and the LCS results pass the continuing calibration verification acceptance criteria. | Persuasive | Language modified to include "initial" but no other changes (not entirely as suggested)
| Negative with comment | 1.7.2 (f) | 1. It is not TNI's right or obligation to determine what data is "valid" for individual laboratory, industry, or regulatory use. This section implies that if a continuing calibration verification fails during an analysis the data is not reported unless it meets the special conditions in 1.7.2.f.iii.a and 1.7.2.f.iii.b. This section should be modified to allow for reporting the data with qualification unless the special conditions in these 2 sections are met. Let the end-user determine if the data is "valid" for their purposes. 2. Most instruments are automated in today's society and these sections do not give options for the course of action to take once the analysts determines what occurred during an automated analytical run where the samples and other QC have already been run before the analyst was aware of the failing continuing calibration verification. 1.7.2.f.i refers to running a second calibration verification sample immediately, but that is not realistic for automated analyses. And especially for a "missed autosampler injection", at least one sample would be in the process of being analyzed prior to the analyst discovering that an injection was missed. | Persuasive | Language modified.
| Negative with comment | 1.7.2 (f) (ii) | The allowance to run two passing CCVs after a failed CCV and then continuing with analysis has been removed. This should be added back in. | Non-Persuasive | The allowance for running two CCVs after a failing CCV was controversial and cannot be added at this stage of the standard. However, we note that it is OK to run a passing CCV after corrective action. The laboratory SOP should define what constitutes corrective action for individual methods.
| Negative with comment | 1.7.2 (f) (iii) | The current language reads: "if samples are analyzed using a system on which the calibration has not been verified, the results shall be qualified. Data associated with an unacceptable calibration verification may be reported under the following special conditions, unless" prohibited by the client, a regulatory program or regulation:, The language needs to be more clear that in such circumstances, if data are to be reported they must also be qualified. Proposed revised language: "if samples are analyzed using a system on which the calibration has not been verified, the results shall be qualified. Data associated with an unacceptable calibration verification may be reported with qualifiers under the following special conditions, unless prohibited by the client, a regulatory program or regulation:, " | Persuasive | Language modified